DIRECTIONS FOR RUNNING THE PLATE POURING APPARATUS

NOTE: You must use Falcon Plate #1029 in this instrument.  Other brands will stick in the carousel.

- Starting from the position labeled “stack 40 dishes maximum”, load the Falcon Plates into the carousel by placing 5 plates at a time into the holders.  Start at the top of the holder and spread the rods apart then move the plates to the bottom.  Repeat until 40 plates are loaded.

- If you need to calibrate, change the dispense volume, cool time or speed refer to the manual.

- Using aseptic technique, carefully unwrap the sterilized tubing and canula.  Place the tubing into the flask in the media and cover with foil.  Place the canula into its holder.

- Push the Auto/Manual button to select the manual mode.

- Press the individual plate carousel button twice to cause a plate to drop and rotate into the position directly below the canula.

- Depress the start button long enough to fill the tubing with media and to dispense any bubbles in the line into the plate.  Hit Stop once this is done.

- Using the manual mode, move the individual plate carousel to clear the used plate.

- Depress the Auto/ Manual button to select the Auto Mode.

- Press Start.

*Plates should drop and fill automatically.

*Be sure to monitor the plate pourer and be prepared to stop it if necessary.

*Clean any spills of agar immediately and thoroughly or the equipment will jam.

- Stop the plate pourer when the media is nearly all dispensed.

- Immediately flush the tubing and canula with hot water.  Rewrap and autoclave tubing for the next user.